CDRL A004

Product Drawings/Models and

Associated Lists for

Advanced Modular Manikin Project

Phase II Program

Contract # W81XWH-14-C-0101

Date: February 3, 2020

Revision: Rev - 2

Unclassified

PRINCIPAL INVESTIGATOR: ROBERT M. SWEET, MD, FACS

REGENTS OF THE UNIVERSITY OF MINNESOTA

OFFICE OF SPONSORED PROJECTS

200 OAK ST SE

MINNEAPOLIS MN 55455-2009

This document is unclassified, and APPROVED FOR PUBLIC RELEASE; DISTRIBUTION IS UNLIMITED

Revision History

| Date of Revision | Description of Changes |
| --- | --- |
| 24 Dec 2018 | Original Draft Submitted |
| 25 Sep 2019 | Document Submitted |
| 03 Feb 2020 | Distribution Updated |

Distribution Statement

This document was created under contract funding W81XWH-14-C-0101 from:

DEPARTMENT OF THE ARMY

US ARMY MEDICAL RESEARCH ACQUISITION ACTIVITY

820 CHANDLER STREET

FORT DETRICK MD 21701-5014

It contains no proprietary information, trade secrets, copy righted material or classified information. This document is unclassified, and APPROVED FOR PUBLIC RELEASE; DISTRIBUTION IS UNLIMITED

Table of Contents

1	Scope	4

1.1	Identification	4

1.2	System Overview	5

1.3	Document Overview	5

2	Referenced Documents	5

2.1	Industry Documents	5

3	Reference Design	6

3.1	The drawing reference document	6

3.2	Commercial Item Drawings	9

3.3	Embedded Electronics Design Information	38

3.3.1	AMM Development Kit	38

3.3.2	AMM Application Board	70

3.3.3	AMM Network Manager	88

Figure 1: Functional Overview of AMM Platform	4

Table 1: AMM Network Manager power capabilities.	97

# Scope

This document defines the standards for 1.0 release of the Advanced Modular Manikin (AMM) platform and its formal deliverables. The formal deliverables consist of the platform specification, an open source* Reference Implementation (RI) of the Computer Software Configuration Items (CSCIs), a reference implementation of the Universal Segment Connector (USC) and other hardware defined by the Hardware Configuration Items (HWCIs), the data models that ensure interoperability between the core and modules, and the documents that describe their design, operation, and extensibility through the addition of AMM Modules. Modules are defined as independent building blocks that provide incremental capabilities to the core or provide training opportunities for different medical and trauma related conditions. The focus of this specification is on the platform, a much broader definition than a physical manikin, as illustrated in Figure 1, and on how it can be extended by medical simulation developers by adding:

- Modules that provide incremental capabilities to the core, including authoring tools, after action review tools, different physiology engines.

- Modules that add training opportunities, including IV/IO arms, intubation heads, laparotomy abdomens, virtual stethoscopes. These can be physical, virtual, or hybrid part task trainers.

Figure 1: Functional Overview of AMM Platform

## Identification

This CDRL Item A004 of Contract # W81XWH-14-C-0101, Phase II provides AMM product drawings/models and associated lists.

This CDRL is formatted to the requirements of Data Item Description Number DI-Sess-81000E as required. This DID is applicable to acquisitions of military systems, equipment, and components. It is intended for acquiring drawings/models and Associated Lists at the end of the System Development and Demonstration Phase and during subsequent phases of the DoD materiel life cycle.

## System Overview

The AMM platform is a modular, distributed, interoperable system that enables physical, virtual, augmented and hybrid modules to work together as an integrated system.  The traditional “core”, i.e. computer and state engine, can be in any one of the traditional manikin segments, i.e. torso, leg etc., or external to the human form, as it would be if the system is only running a virtual instance or if the targeted scenario, i.e. patient case, does not allow them to be internal due to the set of interventions that have to be performed on the body.  The platform is architected as a system of systems that allow modules to function either as part of an integrated, whole body simulation or as autonomous part task trainers.

The published AMM standards guide the development and integration of AMM compatible modules.  The reference designs provided for the final demo including electronics and central supplies were created to demonstrate the operation of the platform and are published as a developer’s tool kit with sources to acquire them from.

The developers of the platform have agreed to publish the AMM platform under the following open source licensing option:

* Creative Commons Attribution 4.0 International (CC BY 4.0) https://creativecommons.org/licenses/by/4.0/deed.ast.

Share — copy and redistribute the material in any medium or format

Adapt — remix, transform, and build upon the material for any purpose, even commercially.

The licensor cannot revoke these freedoms as long as you follow the license terms.

This document does not cover modules that were created under separate funding and by other entities to demonstrate the functionality of the AMM Platform under separate funding and are not part of the Open Source agreement.

## Document Overview

The outline and subject matter content of CDRL A004 are based on DI-SESS-81000E as required by the contract. The DID has been tailored as appropriate.

# Referenced Documents

## Industry Documents

| Doc. No. | Title |
| --- | --- |
| MLL-STD-31000 | Technical Data Package |
| W81XWH-14-C-0101 | AMM Phase II Contract, DOD |
| DI-SESS-81000E | Data Item Description |

# Reference Design

The reference system design CAD is in SolidWorks.  The drawing reference document (Dwg. No. 100100) includes the system schematic design documents and an indented bill of materials for the mechanical assembly. The drawings/models and associated lists are contractor format.

Drawing reference items with 6-digit numerical part numbers are contractor assemblies and manufactured parts. All commercial items are referenced with drawing number indicating OEM in short form and OEM part number. Commercial items are further documented in CDRL A003.

## The drawing reference document

Starting on following page.

## Contractor Assemblies and Manufactured Item Drawings

Starting on following page.

## Embedded Electronics Design Information

Design information for the custom designed AMM embedded electronics items are given in schematic form in contractor format.

### AMM Development Kit

#### AMM Development Kit Input / Output Specification

GPIO Connectors J16 (Left) , J17 (Right)

Power Header J28

GPIO Connectors J16 (Left) , J17 (Right)

Power Header J28

Power Header J28

| Pin Number | Pin Name | Description | Max Current |
| --- | --- | --- | --- |
| 1 | 5V | Standby Power (Low current only) |  |
| 2 | 12V_DC_JACK_IN | 12 V Input Sense (Drive at 12 V when supplying current from the application board) |  |
| 3 | MB_VREG_3P3 | 3.3 V Main Rail | 8 A |
| 4 | MB_VREG_3P3 | 3.3 V Main Rail | 8 A |
| 5 | GND | Common Ground |  |
| 6 | GND | Common Ground |  |
| 7 | MB_VREG_5P0 | 5 V Main Rail | 8 A |
| 8 | MB_VREG_5P0 | 5 V Main Rail | 8 A |
| 9 | GND | Common Ground |  |
| 10 | GND | Common Ground |  |
| 11 | GND | Common Ground |  |
| 12 | GND | Common Ground |  |
| 13 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) | 8 A |
| 14 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) | 8 A |
| 15 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) | 8 A |
| 16 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) | 8 A |

| J16 Pin Number | Pin Name | Low Pass Filter | Digital Input | Digital Output | PWM Output | Analog Input | Analog Output (DAC) | Serial Interface | Other |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | PTD6_RC |  | PTD6 | PTD6 |  | ADC0_SE7b |  | SPI1_SOUT, UART0_RX | FTM0_CH6, FTM0_FLT0 |
| 2 | GND |  |  |  |  |  |  |  |  |
| 3 | PTD4_RC |  | PTD4 | PTD4 |  |  |  | SPI1_PSC0, UART0_RTS_b | FTM0_CH4 |
| 4 | PTD5_RC |  | PTD5 | PTD5 |  | ADC0_SE6b |  | SPI1_SCK, UART0_CTS_b, UART0_COL_b | FTM0_CH5 |
| 5 | PTC8_RC |  | PTC8 | PTC8 |  | CMP0_IN0 |  | I2S0_RX_BCLK, I2S0_MCLK | PDB0_EXTRG |
| 6 | PTC9_RC |  | PTC9 | PTC9 |  | ADC1_SE5b, CMP0_IN3 |  | I2S0_RX_BCLK | FTM2_FLT0, FTM3_CH5 |
| 7 | PTA4_RC |  | PTA4 | PTA4 |  |  |  |  | FTM0_CH1, TSI0_CH5 |
| 8 | PTC0_RC |  | PTC0 | PTC0 |  | ADC0_SE14 |  | UART1_RTS_b, I2S0_TXD1 | TSI0_CH13, PDB0_EXTRG |
| 9 | PTB1_RC |  | PTB1 | PTB1 | TPM1_CH1 | ADC0_SE9, ADC1_SE9 |  |  | FTM1_CH1, FTM1_QD_PHB |
| 10 | PTC1_RC |  | PTC1 | PTC1 |  | ADC0_SE15 |  | UART1_RTS_b, I2S0_TXD0 | TSI0_CH14, FTM0_CH0 |
| 11 | GND |  |  |  |  |  |  |  |  |
| 12 | PTA8_RC |  | PTA8 | PTA8 | TPM1_CH0 | ADC0_SE11 |  |  | FTM1_CH0, FTM1_QD_PHA |
| 13 | PTB4_RC |  | PTB4 | PTB4 |  | ADC1_SE10 |  |  | FTM1_FLT0 |
| 14 | PTB3_RC |  | PTB3 | PTB3 |  | ADC0_SE13 |  | I2C0_SDA, UART0_CTS_b, UART0_COL_b | FTM0_FLT0 |
| 15 | ADC1_SE15_RC |  | PTB11 | PTB11 |  | ADC1_SE15 |  | SPI1_SCK, UART3_TX | FTM0_FLT2 |
| 16 | ADC1_SE14_RC |  | PTB10 | PTB10 |  | ADC1_SE14 |  | SPI1_PCS0, UART3_RX | FTM0_FLT1 |
| 17 | PTB6_RC |  | PTB6 | PTB6 |  | ADC1_SE12 |  |  |  |
| 18 | PTB7_RC |  | PTB7 | PTB7 |  | ADC1_SE13 |  |  |  |
| 19 | PTB2_RC |  | PTB2 | PTB2 |  | ADC0_SE12 |  | I2C0_SCL, UART0_RTS_b | FTM0_FLT3 |
| 20 | PTB5_RC |  | PTB5 | PTB5 |  | ADC1_SE11 |  |  | FTM2_FLT0 |
| 21 | PTB0_RC |  | PTB0 | PTB0 | TPM1_CH0 | ADC0_SE8, ADC1_SE8 |  | I2C0_SCL | FTM1_CH0, FTM1_QD_PHA |
| 22 | PTA11_RC |  | PTA11 | PTA11 | TPM2_CH1 |  |  | I2C2_SDA | FTM2_CH1, FTM2_QD_PHB |
| 23 | PTA13_RC |  | PTA13 | PTA13 | TPM1_CH1 | CMP2_IN1 |  | CAN0_RX, I2C2_SDA, I2S0_TX_FS | FTM1_QD_PHB |
| 24 | PTA14_RC |  | PTA14 | PTA14 |  |  |  | UART0_TX, I2C2_SCL, I2S0_RX_BCLK, I2S0_TXD1 |  |
| 25 | PTA17_RC |  | PTA17 | PTA17 |  | ADC1_SE17 |  | UART0_RTS_b, I2S0_MCLK |  |
| 26 | PTA15_RC |  | PTA15 | PTA15 |  | CMP3_IN1 |  | UART0_RX, I2S0_RXD0 |  |
| 27 | PTA10_RC |  | PTA10 | PTA10 | TPM2_CH0 |  |  |  | FTM2_QD_PHA |
| 28 | GND |  |  |  |  |  |  |  |  |
| 29 | ADC0_SE18_RC |  | PTE25 | PTE25 |  | ADC0_SE18 |  | CAN1_RX, UART4_RX, I2C0_SDA |  |
| 30 | PTA12_RC |  | PTA12 | PTA12 | TPM1_CH0 | CMP2_IN0 |  | CAN0_TX, I2C2_SCL, I2S0_TXD0 | FTM1_CH0, FTM1_QD_PHA |
| 31 | ADC1_SE18_RC |  |  |  |  | ADC1_SE18, CMP1_IN5, CMP0_IN5 |  |  | VREF_OUT |
| 32 | PTA7_RC |  | PTA7 | PTA7 |  | ADC0_SE10 |  |  | FTM0_CH4 |
| 33 | ADC1_SE16_RC |  |  |  |  | ADC1_SE16, CMP2_IN2, ADC0_SE22 |  |  |  |
| 34 | ADC0_SE17_RC |  | PTE24 | PTE24 |  | ADC0_SE17 |  | CAN1_TX, UART4_TX, I2C0_SCL |  |
| 35 | ADC1_DP0_RC |  |  |  |  | ADC1_DP0, ADC0_DP3 |  |  |  |
| 36 | ADC0_SE16_RC |  |  |  |  | ADC0_SE16, CMP1_IN2, ADC0_SE21 |  |  |  |
| 37 | ADC0_DM0_RC |  |  |  |  | ADC0_DM0, ADC1_DM3 |  |  |  |
| 38 | ADC1_DM0_RC |  |  |  |  | ADC1_DM0, ADC0_DM3 |  |  |  |
| 39 | GND |  |  |  |  |  |  |  |  |
| 40 | RST_TGTMCU_B |  |  |  |  |  |  |  | RESET K66 MCU |

| J17 Pin Number | Pin Name | Low Pass Filter | Digital Input | Digital Output | PWM Output | Analog Input | Analog Output (DAC) | Serial Interface | Other |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | PTD11 |  | PTD11 | PTD11 |  |  |  | SPI2_PCS0, LPUART0_CTS_b | SDHC0_CLKIN |
| 2 | GND |  |  |  |  |  |  |  |  |
| 3 | PTD12 |  | PTD12 | PTD12 |  |  |  | SPI2_SCK | FTM3_FLT0 |
| 4 | PTC14 |  | PTC14 | PTC14 |  |  |  | SPI2_SIN |  |
| 5 | I2C0_SDA |  | PTD9 | PTD9 |  |  |  | I2C0_SDA, LPUART0_TX |  |
| 6 | FTM0_CH3 |  | PTC4 | PTC4 |  |  |  | UART1_TX | LLWU_P8, FTM0_CH3, CMP1_OUT |
| 7 | I2C0_SCL |  | PTD8 | PTD8 |  |  |  | I2C0_SCL, LPUART0_RX | LLWU_P24 |
| 8 | I2C1_SDA |  | PTC11 | PTC11 |  | ADC1_SE7b |  | I2C1_SDA, I2S0_RXD1 | LLWU_P11, FTM3_CH7 |
| 9 | PTD7 |  | PTD7 | PTD7 |  |  |  | UART0_TX, SPI1_SIN | FTM0_FLT1, CMT_IRO, FTM0_CH7 |
| 10 | I2C1_SCL |  | PTC10 | PTC10 |  | ADC1_SE6b |  | I2C1_SCL, I2S0_RX_FS | FTM3_CH6 |
| 11 | PTC15 |  | PTC15 | PTC15 |  |  |  | UART4_TX |  |
| 12 | FTM0_CH2 |  | PTC3 | PTC3 |  |  |  | UART1_RX, I2S0_TX_BCLK | CMP1_IN1, LLWU_P7, FTM0_CH2, CLKOUT |
| 13 | GND |  |  |  |  |  |  |  |  |
| 14 | FTM0_CH1 |  | PTC2 | PTC2 |  | ADC0_SE4b |  | UART1_CTS_b, I2S0_TX_FS | CMP1_IN0, TSI0_CH15, FTM0_CH1 |
| 15 | PTB21 |  | PTB21 | PTB21 |  |  |  | SPI2_SCK | CMP1_OUT |
| 16 | PTB23 |  | PTB23 | PTB23 |  |  |  | SPI2_SIN | CMP3_OUT |
| 17 | PTB20 |  | PTB20 | PTB20 |  |  |  | SPI2_PCS0 | CMP0_OUT |
| 18 | PTB22 |  | PTB22 | PTB22 |  |  |  | SPI2_SOUT | CMP2_OUT |
| 19 | PTB18 |  | PTB18 | PTB18 | TPM2_CH0 |  |  | CAN0_TX, I2S0_TX_BCLK | TSI0_CH11, FTM2_CH0, FTM2_QD_PHA |
| 20 | PTB19 |  | PTB19 | PTB19 | TPM2_CH1 |  |  | CAN0_RX, I2S0_TX_FS | FTM2_CH1, FTM2_QD_PHB, TSI0_CH12 |
| 21 | SDHC0_DCLK |  | PTE2 | PTE2 |  | ADC1_SE6a |  | SPI1_SCK, UART1_CTS_b | LLWU_P1, SDHC0_DCLK, TRACE_D2 |
| 22 | SD_CARD_DETECT |  | PTD10 | PTD10 |  |  |  | LPUART0_RTS_b |  |
| 23 | SDHC0_D1 |  | PTE0 | PTE0 |  | ADC1_SE4a |  | SPI1_PCS1, UART1_TX, I2C1_SDA | SDHC0_D1, TRACE_CLKOUT, RTC_CLKOUT |
| 24 | SDHC0_D0 |  | PTE1 | PTE1 |  | ADC1_SE5a |  | SPI1_SOUT, UART1_RX, I2C1_SCL, SPI1_SIN | LLWU_P0, SDHC0_D0, TRACE_D3 |
| 25 | SDHC0_D3 |  | PTE4 | PTE4 |  |  |  | SPI1_PCS0, UART3_TX | SDHC0_D3, TRACE_D0 |
| 26 | SDHC0_CMD |  | PTE3 | PTE3 |  | ADC1_SE7a |  | SPI1_SIN, UART1_RTS_b, SPI1_SOUT | SDHC0_CMD, TRACE_D1 |
| 27 | PTE6 |  | PTE6 | PTE6 |  |  |  | SPI1_PCS3, UART3_CTS_b, I2S0_MCLK | FTM3_CH1 |
| 28 | SDHC0_D2 |  | PTE5 | PTE5 |  |  |  | SPI1_PCS2, UART3_RX, SDHC0_D2 | FTM3_CH0 |
| 29 | PTE8 |  | PTE8 | PTE8 |  |  |  | I2S0_RXD1, I2S0_RX_FS, LPUART0_TX | FTM3_CH3 |
| 30 | PTE7 |  | PTE7 | PTE7 |  |  |  | UART3_RTS_b, I2S0_RXD0 | FTM3_CH2 |
| 31 | PTE11 |  | PTE11 | PTE11 |  |  |  | I2S0_TX_FS, LPUART0_RTS_b | FTM3_CH6 |
| 32 | PTE12 |  | PTE12 | PTE12 |  |  |  | I2S0_TX_BCLK | FTM3_CH7 |
| 33 | CAN0_H |  | PTC17 | PTC17 |  |  |  | CAN1_TX, UART3_TX |  |
| 34 | CAN0_L |  | PTC16 | PTC16 |  |  |  | CAN1_RX, UART3_RX |  |
| 35 | DAC0_OUT_RC |  |  |  |  | ADC0_SE23 |  | - | CMP1_IN3 |
| 36 | DAC1_OUT_RC |  |  |  |  | ADC1_SE23 |  | - | CMP0_IN4, CMP2_IN3 |
| 37 | PTE26 |  | PTE26 | PTE26 |  |  |  | UART4_CTS_b | RTC_CLKOUT |
| 38 | PTE28 |  | PTE28 | PTE28 |  |  |  | - |  |
| 39 | GND |  |  |  |  |  |  |  |  |
| 40 | RESERVED FOR INTERNAL USE |  |  |  |  |  |  |  |  |

#### AMM Development Kit Schematic

### AMM Application Board

#### I/O specifications

| Capabilities |  |  |  |
| --- | --- | --- | --- |
| Function | Quantity | Description | Notes |
| Solenoids | 16 | 12 V solenoid drivers | 2 are PWM capable |
| Pressure sensors | 8 | 0-103.4 kPa (0-15 PSI) | 16-bit resolution |
| Motor Driver | 1 | Three phase brushless DC motor driver | 12 V max output |
| Low Flow Sensor | 1 | 0 mL/min - 240 mL/min | I2C interface |
| High Flow Sensor | 1 | 100 mL/min - 2500 mL/min | 22000 pulses/L |
| GPIO | 7 | General Purpose Input and Output pins | 3.3 V max input, 25 mA per pin, 100 mA total current draw |
| Analogue Inputs | 2 | 16-bit Analogue to Digital Converters | 3.3 V max input |
| Serial Interfaces | 3 | I2C, SPI, and UART | 3.3 V max input |

Power Header J2 (Power From Motherboard)

| Pin Number | Pin Name | Description | Notes |
| --- | --- | --- | --- |
| 1 | 5V | Standby Power (Low current only) |  |
| 2 | 12V_DC_JACK_IN | 12 V Input Sense (Drive at 12 V when supplying current from the application board) |  |
| 3 | MB_VREG_3P3 | 3.3 V Main Rail | 5 A max |
| 4 | MB_VREG_3P3 | 3.3 V Main Rail |  |
| 5 | GND | Common Ground |  |
| 6 | GND | Common Ground |  |
| 7 | MB_VREG_5P0 | 5 V Main Rail | 5 A max |
| 8 | MB_VREG_5P0 | 5 V Main Rail |  |
| 9 | GND | Common Ground |  |
| 10 | GND | Common Ground |  |
| 11 | GND | Common Ground |  |
| 12 | GND | Common Ground |  |
| 13 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) | 5 A max |
| 14 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) |  |
| 15 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) |  |
| 16 | VCC_12V | 12 V Main Rail, Bi-directional (High Current) |  |

GPIO J16 (From Motherboard)

| J16 Pin Number | Pin Name | Generic Application Board Pin Function | Usage |
| --- | --- | --- | --- |
| 1 | PTD6_RC | Soleonid 3 Control | Set high to activate solenoid |
| 2 | GND | Common Ground |  |
| 3 | PTD4_RC | Solenoid 1 Control | Set high to activate solenoid |
| 4 | PTD5_RC | Solenoid 2 Control | Set high to activate solenoid |
| 5 | PTC8_RC | Solenoid 12 Control | Set high to activate solenoid |
| 6 | PTC9_RC | High volume flow sensor | Count pulses to determine flow rate |
| 7 | PTA4_RC | Solenoid 13 Control | Set high to activate solenoid |
| 8 | PTC0_RC | GPIO 7 | 3.3 V max |
| 9 | PTB1_RC | Motor Enable | Set high to activate motor |
| 10 | PTC1_RC | Solenoid 10 Control | Set high to activate solenoid |
| 11 | GND | Common Ground |  |
| 12 | PTA8_RC | Solenoid 11 Control | Set high to activate solenoid |
| 13 | PTB4_RC | Motor Speed Feedback (Analogue) | Read voltage level to determine motor speed |
| 14 | PTB3_RC | Motor IO | General purpose motor IO |
| 15 | ADC1_SE15_RC | Analogue Input 1 | Read voltage level (3.3 V max) |
| 16 | ADC1_SE14_RC | Pressure Sensor 4 (Analogue) | Read voltage level to determine pressure |
| 17 | PTB6_RC | Pressure Sensor 2 (Analogue) | Read voltage level to determine pressure |
| 18 | PTB7_RC | Pressure Sensor 3 (Analogue) | Read voltage level to determine pressure |
| 19 | PTB2_RC | Motor Status Feedback |  |
| 20 | PTB5_RC | Pressure Sensor 1 (Analogue) | Read voltage level to determine pressure |
| 21 | PTB0_RC | Motor PWM Speed Control | Modulate pulse width to set motor speed |
| 22 | PTA11_RC | Solenoid 6 Control | Set high to activate solenoid (PWM capable) |
| 23 | PTA13_RC | J15 I2C Header SDA | I2C2, 3.3 V max |
| 24 | PTA14_RC | J3 UART TX | UART transmit (3.3 V) |
| 25 | PTA17_RC | Solenoid 14 Control | Set high to activate solenoid |
| 26 | PTA15_RC | J3 UART RX | UART receive (3.3 V max) |
| 27 | PTA10_RC | Solenoid 5 Control | Set high to activate solenoid (PWM capable) |
| 28 | GND | Common Ground |  |
| 29 | ADC0_SE18_RC | Pressure Sensor 5 (Analogue) | Read voltage level to determine pressure |
| 30 | PTA12_RC | J15 I2C Header SCL | I2C2, 3.3 V max |
| 31 | ADC1_SE18_RC | Pressure Sensor 6 (Analogue) | Read voltage level to determine pressure |
| 32 | PTA7_RC | Solenoid 9 Control | Set high to activate solenoid |
| 33 | ADC1_SE16_RC | Analogue Input 2 | Read voltage level (3.3 V max) |
| 34 | ADC0_SE17_RC | Solenoid 7 Control | Set high to activate solenoid |
| 35 | ADC1_DP0_RC | Pressure Sensor 7 (Analogue) | Read voltage level to determine pressure |
| 36 | ADC0_SE16_RC | N/A |  |
| 37 | ADC0_DM0_RC | Pressure Sensor 8 (Analogue) | Read voltage level to determine pressure |
| 38 | ADC1_DM0_RC | N/A |  |
| 39 | GND | Common Ground |  |
| 40 | RST_TGTMCU_B | N/A |  |

GPIO J17 (From Motherboard)

| J17 Pin Number | Pin Name | Generic Application Board Pin Function | Usage |
| --- | --- | --- | --- |
| 1 | PTD11 | Solenoid 15 Control | Set high to activate solenoid |
| 2 | GND | Common Ground |  |
| 3 | PTD12 | Solenoid 16 Control | Set high to activate solenoid |
| 4 | PTC14 | N/A |  |
| 5 | I2C0_SDA | N/A |  |
| 6 | FTM0_CH3 | N/A |  |
| 7 | I2C0_SCL | N/A |  |
| 8 | I2C1_SDA | Low Flow Sensor I2C SDA | I2C1, 3.3 V max |
| 9 | PTD7 | Solenoid 4 Control | Set high to activate solenoid |
| 10 | I2C1_SCL | Low Flow Sensor I2C SCL | I2C1, 3.3 V max |
| 11 | PTC15 | N/A |  |
| 12 | FTM0_CH2 | N/A |  |
| 13 | GND | Common Ground |  |
| 14 | FTM0_CH1 | N/A |  |
| 15 | PTB21 | GPIO 4 | 3.3 V max |
| 16 | PTB23 | GPIO 6 | 3.3 V max |
| 17 | PTB20 | GPIO 3 | 3.3 V max |
| 18 | PTB22 | GPIO 5 | 3.3 V max |
| 19 | PTB18 | GPIO 1 | 3.3 V max |
| 20 | PTB19 | GPIO 2 | 3.3 V max |
| 21 | SDHC0_DCLK | SPI SCLK | SPI1, 3.3 V max |
| 22 | SD_CARD_DETECT | N/A |  |
| 23 | SDHC0_D1 | N/A |  |
| 24 | SDHC0_D0 | SPI MOSI | SPI1, 3.3 V max |
| 25 | SDHC0_D3 | SPI SS | SPI1, 3.3 V max |
| 26 | SDHC0_CMD | SPI MISO | SPI1, 3.3 V max |
| 27 | PTE6 | N/A |  |
| 28 | SDHC0_D2 | N/A |  |
| 29 | PTE8 | N/A |  |
| 30 | PTE7 | N/A |  |
| 31 | PTE11 | N/A |  |
| 32 | PTE12 | Solenoid 8 Control | Set high to activate solenoid |
| 33 | CAN0_H | N/A |  |
| 34 | CAN0_L | N/A |  |
| 35 | DAC0_OUT_RC | N/A |  |
| 36 | DAC1_OUT_RC | Motor Speed Input (Analogue) | Set an analogue voltage to control motor speed |
| 37 | PTE26 | N/A |  |
| 38 | PTE28 | N/A |  |
| 39 | GND | Common Ground |  |
| 40 | RESERVED FOR INTERNAL USE | N/A |  |

Generic Application Board Solenoid Control

| Descriptor | Pin | Function | Motherboard MCU Pin Name | Notes |
| --- | --- | --- | --- | --- |
| J4 | 8 | Solenoid 1 GND (N-FET Controlled) | PTD4_RC | 500 mA max* |
| J4 | 7 | Solenoid 2 GND (N-FET Controlled) | PTD5_RC | 500 mA max* |
| J4 | 6 | Solenoid 3 GND (N-FET Controlled) | PTD6_RC | 500 mA max* |
| J4 | 5 | Solenoid 4 GND (N-FET Controlled) | PTD7 | 500 mA max* |
| J4 | 4 | Solenoid 5 GND (N-FET Controlled) | PTA10_RC | 500 mA max*, PWM Capable |
| J4 | 3 | Solenoid 6 GND (N-FET Controlled) | PTA11_RC | 500 mA max*, PWM Capable |
| J4 | 2 | Solenoid 7 GND (N-FET Controlled) | ADC0_SE17_RC | 500 mA max* |
| J4 | 1 | Solenoid 8 GND (N-FET Controlled) | PTE12 | 500 mA max* |
| J5 | 1-8 | Solenoids 1-8 +12V Supply | N/A |  |
| J6 | 8 | Solenoid 9 GND (N-FET Controlled) | PTA7_RC | 500 mA max* |
| J6 | 7 | Solenoid 10 GND (N-FET Controlled) | PTC1_RC | 500 mA max* |
| J6 | 6 | Solenoid 11 GND (N-FET Controlled) | PTA8_RC | 500 mA max* |
| J6 | 5 | Solenoid 12 GND (N-FET Controlled) | PTC8_RC | 500 mA max* |
| J6 | 4 | Solenoid 13 GND (N-FET Controlled) | PTA4_RC | 500 mA max* |
| J6 | 3 | Solenoid 14 GND (N-FET Controlled) | PTA17_RC | 500 mA max* |
| J6 | 2 | Solenoid 15 GND (N-FET Controlled) | PTD11 | 500 mA max* |
| J6 | 1 | Solenoid 16 GND (N-FET Controlled) | PTD12 | 500 mA max* |
| J7 | 1-8 | Solenoids 9-16 +12V Supply | N/A |  |
| Note* | Consult the datasheet for the IC for recommended current per pin when using multiple pins at varying duty cyles. |  |  |  |
|  | TBD62083A Datasheet |  |  |  |

Generic Application Board Pressure Sensor

| Descriptor | Pressure Sensor Name | Motherboard MCU Pin Name | Notes |  |  |
| --- | --- | --- | --- | --- | --- |
| U3 | Pressure 1 | PTB5_RC | To calculate pressure, read the 12-bit ADC value, and use this equation: Pressure (PSI) = VALUE * (3/10280*16) - (15/8) |  |  |
| U4 | Pressure 2 | PTB6_RC |  |  |  |
| U5 | Pressure 3 | PTB7_RC |  |  |  |
| U6 | Pressure 4 | ADC1_SE14_RC |  |  |  |
| U8 | Pressure 5 | ADC0_SE18_RC |  |  |  |
| U9 | Pressure 6 | ADC1_SE18_RC |  |  |  |
| U10 | Pressure 7 | ADC1_DP0_RC |  |  |  |
| U11 | Pressure 8 | ADC0_DM0_RC |  |  |  |

| Sensirion SLQ-QT500 Low Flow Sensor (0 mL/min - 240 mL/min) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Descriptor | Pin | Description | Motherboard MCU Pin Name | Notes |  |
| J8 | 1 | FLOW_SDA (I2C) | I2C1_SDA | 3.3 V max |  |
| J8 | 2 | GND | N/A |  |  |
| J8 | 3 | 3.3 V Supply | N/A | 5 A max |  |
| J8 | 4 | FLOW_SCL (I2C) | I2C1_SCL | 3.3 V max |  |
| Omega FTB421 High Flow Sensor (100 mL/min - 2500 mL/min) |  |  |  |  |  |
| Descriptor | Pin | Description | Motherboard MCU Pin Name | Notes |  |
| J8 | 6 | 5 V Supply | N/A | 5 A max |  |
| J8 | 7 | Flow Pulse Counter | PTC9_RC | 22000 pulses/L |  |
| J8 | 8 | GND | N/A |  |  |

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn | Notes |
| --- | --- | --- | --- | --- | --- |
| J15 | 1 | 3.3 V Supply | I2C | N/A | 5 A max |
| J15 | 2 | SCL | I2C | PTA12_RC | 3.3 V max |
| J15 | 3 | SDA | I2C | PTA13_RC | 3.3 V max |
| J15 | 4 | GND | I2C | N/A |  |

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn | Notes |
| --- | --- | --- | --- | --- | --- |
| J1 | 1 | 3.3 V Supply | SPI | N/A | 5 A max |
| J1 | 2 | GND | SPI | N/A |  |
| J14 | 1 | SCLK | SPI | SDHC0_DCLK | 3.3 V max |
| J14 | 2 | MISO | SPI | SDHC0_D0 | 3.3 V max |
| J14 | 3 | MOSI | SPI | SDHC0_CMD | 3.3 V max |
| J14 | 4 | SS | SPI | SDHC0_D3 | 3.3 V max |

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn | Notes |
| --- | --- | --- | --- | --- | --- |
| J3 | 1 | 3.3 V Supply | UART | N/A | 5 A max |
| J3 | 2 | RX | UART | PTA15_RC | 3.3 V max |
| J3 | 3 | TX | UART | PTA14_RC | 3.3 V max |
| J3 | 4 | GND | UART | N/A |  |

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn | Notes |
| --- | --- | --- | --- | --- | --- |
| J9 | 1 | GPIO 1 | GPIO | PTB18 | 3.3 V max, PWM Capable |
| J9 | 2 | GPIO 2 | GPIO | PTB19 | 3.3 V max, PWM Capable |
| J9 | 3 | GPIO 3 | GPIO | PTB20 | 3.3 V max |
| J9 | 4 | GPIO 4 | GPIO | PTB21 | 3.3 V max |
| J9 | 5 | GPIO 5 | GPIO | PTB22 | 3.3 V max |
| J9 | 6 | GPIO 6 | GPIO | PTB23 | 3.3 V max |
| J9 | 7 | GPIO 7 | GPIO | PTC0_RC | 3.3 V max, Analogue Input Capable |
| J9 | 8 | GND | GPIO | N/A |  |

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn | Notes |
| --- | --- | --- | --- | --- | --- |
| J22 | 1 | 3.3 V Supply | Analogue In | N/A | 5 A max |
| J22 | 2 | Analogue Input 1 | Analogue In | ADC1_SE15_RC | 3.3 V max |
| J22 | 3 | Analogue Input 2 | Analogue In | ADC1_SE16_RC | 3.3 V max |
| J22 | 4 | GND | Analogue In | N/A |  |

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn | Notes |
| --- | --- | --- | --- | --- | --- |
| J20 | 1 | 3.3 V Supply | Power Out | N/A | 5 A max |
| J20 | 2 | 5 V Supply | Power Out | N/A | 5 A max |
| J20 | 3 | 12 V Supply | Power Out | N/A | 5 A max |
| J20 | 4 | GND | Power Out | N/A |  |

Motor Controller

| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn |
| --- | --- | --- | --- | --- |
| J10 | 1 | Motor Winding 1 | Motor | N/A |
| J10 | 2 | Motor Winding 2 | Motor | N/A |
| J10 | 3 | Motor Winding 3 | Motor | N/A |
| J10 | 4 | Hall Sensor 5 V Power | Motor | N/A |
| J10 | 5 | Hall Sensor 1 | Motor | N/A |
| J10 | 6 | Hall Sensor 2 | Motor | N/A |
| J10 | 7 | Hall Sensor 3 | Motor | N/A |
| J10 | 8 | Hall Sensor GND | Motor | N/A |
| Descriptor | Pin | Function Description | Interface | Motherboard MCU PIn |
| U7 | N/A | MOTOR_ENABLE | Motor | PTB1_RC |
| U7 | N/A | MOTOR_PWM | Motor | PTB0_RC |
| U7 | N/A | MOTOR_STATUS | Motor | PTB2_RC |
| U7 | N/A | MOTOR_IO | Motor | PTB3_RC |
| U7 | N/A | MOTOR_SPEED_IN | Motor | DAC1_OUT_RC |
| U7 | N/A | MOTOR_SPEED_OUT | Motor | PTB4_RC |

1.5.2.2

### AMM Network Manager

#### AMMDK Core Subsystem

The AMMDK Core subsystem implemented in the Network Manager is identical in design to the AMM Development Kit documented in section 3.3.1.

#### Network Switch Subsystem

#### Power over Ethernet Subsystem

Table 1: AMM Network Manager power capabilities.